Correction for background calculation during TADI 2024 campaign and its implication on  CH4 emission rates for reported values by Heidelberg University.

Julia B. Wietzel and Martina Schmidt (Heidelberg University)

Dec 10th 2024

The team from Heidelberg University participated in a one-week controlled blind release campaign from June 24 to 28, 2024, with a mobile measurement setup for ground-based methane emission quantification. The releases were performed on the TADI platform of TotalEnergies in Lacq, France, and organized by Stanford University. The measurements were conducted with a high-precision trace gas analyzer from Licor (LI-7810) and evaluated using a simple Gaussian plume model.

Problem:

During the first analysis and calculation of the CH4 emission rates, within six weeks after the campaign, the background calculation was performed incorrectly due to an error in the programming code.  A constant CH4 offset of 10 ppb (or 0.01 ppm see line in code) was additionally subtracted from the calculated CH4 background The background was incorrectly calculated 10 ppb too low and therefore the CH4 enhancements (peaks) are 10 ppb too high.

Part in code:

bg.baseline <- rfbaseline(x = input_CH4$UTC, y = input_CH4$CH4, NoXP = 500, maxit = c(15,0),

DOT = TRUE, Scale=function(x) mad(x, center=0))

input_CH4$bg <- bg.baseline$fit -0.01 # offset parameter

This effects especially the smaller CH4 concentrations and therefore the lower CH4 release rates as higher CH4 enhancements were calculated due to the wrong background calculation. An example of a time series is shown in Figure 1 a (Release 1 June 24th) with the background fit including the offset error in red and the corrected background fit in blue.

The background was corrected for all release experiments with release rates < 2 kgCH4 h-1 (35% of all releases) as for higher release rates no significant change is expected. A graphical representation of the determined emission rates for the affected range and a table with the corresponding values are shown in Figure 1b and Table1.

For the controlled CH4 release experiments with release rates between 0.5 and 2.0 kgCH4 h-1  the  release rates calculated and reported by Heidelberg  University were estimated 10-30% to high due to the calculation error in the background of the CH4 concentration.

a							 b

Figure 1 a Time series of the measured CH4 concentration during plume crossings for a 45 minute release on June 24th (RE1_1). The background fit affected by the offset error is marked in red and the corrected background fit in blue. b Graphical representation of the change in the affected releases. Bar plot showing the estimated CH4 emission rates with the error in the background calculation and with the corrected background (each as mean and median) compared to the true release rate in blue.

| Release No. | Date | Release rate [kg/h] | Estimated Mean [kg/h] uncorrected | Standard error [kg/h] uncorrected | Estimated Median [kg/h] uncorrected | Standard error (median) [kg/h] uncorrected | Estimated Mean [kg/h] corrected | Standard error [kg/h] corrected | Estimated Median [kg/h] corrected | Standard error (median) [kg/h] corrected |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| RE_3_8 | 2024-06-26 | 0.00 | 0.70 | 0.14 | 0.57 | 0.16 | 0.53 | 0.13 | 0.42 | 0.15 |
| RE_3_5 | 2024-06-26 | 0.01 | 0.64 | 0.18 | 0.44 | 0.06 | 0.21 | 0.03 | 0.23 | 0.06 |
| RE_5_2 | 2024-06-28 | 0.32 | 0.46 | 0.06 | 0.41 | 0.03 | 0.27 | 0.05 | 0.24 | 0.03 |
| RE_4_3 | 2024-06-27 | 0.57 |  |  | 0.80 | 0.26 |  |  | 0.80 | 0.26 |
| RE_1_1 | 2024-06-24 | 0.59 | 0.61 | 0.10 | 0.37 | 0.09 | 0.46 | 0.09 | 0.21 | 0.06 |
| RE_2_1 | 2024-06-25 | 0.77 | 1.42 | 0.18 | 1.54 | 0.24 | 1.28 | 0.18 | 1.41 | 0.23 |
| RE_1_4 | 2024-06-24 | 0.89 | 0.96 | 0.24 | 0.63 | 0.10 | 0.80 | 0.23 | 0.44 | 0.08 |
| RE_3_2 | 2024-06-26 | 1.33 | 1.67 | 0.45 | 1.22 | 0.32 | 1.52 | 0.43 | 1.10 | 0.31 |
| RE_4_8 | 2024-06-27 | 1.44 |  |  | 1.06 | 0.27 |  |  | 1.06 | 0.27 |
| RE_2_8 | 2024-06-25 | 1.48 | 3.03 | 0.46 | 2.63 | 0.43 | 2.73 | 1.43 | 2.17 | 1.15 |
| RE_3_1 | 2024-06-26 | 1.62 | 0.50 | 0.18 | 0.39 | 0.19 | 0.45 | 0.17 | 0.28 | 0.14 |
| RE_2_2 | 2024-06-25 | 1.95 | 2.35 | 0.69 | 1.67 | 0.73 | 2.35 | 0.69 | 1.68 | 0.74 |

Table 1 Overview of the estimated methane emission rates [kgCH4 h-1] of all releases below 2 kgCH4 h-1. The yellow lines are not affected as this was a different measurement method (OTM33-a) with a different type of background calculation. The releases marked red are reported as uncertain, caused by low wind speeds (< 2 m s-1) during the measurements.